## Slide 1
Summary from the users workshop
G.Velev

## Slide 2: Introduction
November 21-22, we had a workshop with the FES user community.
In my opinion the workshop was very successful: 33 registered, ~ 40 people in one time during the meeting
The agenda included
  First day: Introduction to the test facility
  Second day: Input from the user community and discussion
  Link to the workshop: https://indico.fnal.gov/event/57027/
Next slides summarizes the requests from the community
After the workshop we had two meetings:
  First: FNAL team - internal to discuss the requests
  Second: join FNAL and LBNL teams
Next slides summarizes the requests (most of the inputs are from Zach Hartwig,
Nicolai Martovetsky talks and discussion) and output from internal meetings

## Slide 3: Introduction – cont.
Three type of recommendation/requests
  General facility capabilities
  Instrumentation and data acquisition
  User support and resources – to early to be discussed
We graded them by difficulties
Abbreviations:
  P - in the plan
  E – easy to implement
  M - moderate to implement
  D - difficult but not impossible
  VD – very difficult, maybe impossible with this facility
For recommendation/requests which can be implemented and not in the plan - needs discussion how to proceed
Anticryostat
Anticryostat Top Plate
Sample Insert

## Slide 4: Recommendations: General facility capabilities
Cable current: 0 – 100 kA (SC) and 0 – 50 kA (DC)
  [E-M] SC transformer w/ sufficient volt-seconds to last 8 hours if secondary voltages achieved – 30 min continues operation, 8 h DC characterization, SULTAN: when SC transformer is drained takes 4-5 h to charge, to define to the user that if they come at 8 that they have 8 h useful time
  [M-D] DC power supply (up to 50 kA or more) for full quench evolution (dynamics, detection) - 50 kA - sinusoidal from 0 to 50 kA
  Fast SC/DC power supply to maximize IxB cycles per test session (1000’s cycles minimum)
  [P] Ability to reverse sample current for opposite IxB force direction - lifting force from the join, self field is in the different direction, forces, and this coming from misalignment.
Cable temperature: 4 – 50 K
  [M] Supercritical helium (not bath-cooled) with programmable mass flow rates
  [M] Flexibility in cryogenic routing (series vs. parallel in cables, top vs. bottom inlets, etc.) - how to connect the sample – using Tevatron pre-cooler, 2.5 g per second, 8g per leg
  [M-D] Mechanism to cycle sample 4-300-4 K rapidly (1-2 times a day?) for thermal testing - 4.5 K - 100oK, SULTAN warm sample between Friday on Monday, heaters for sample maybe needed

## Slide 5: Recommendations: General facility capabilities
[M] For REBCO cooling should be much stronger, because no or minimal neutron shielding is expected, 10 -15 g LHe flow
External field: 0 – 15 T
  [D] Ability to raise samples and place joints in-field for testing - sample moving – SULTAN has 4 extenders of lengths 40, 160, 330, and 670 mm. They can be stacked to yield various lengths; however, a total of 830 mm should not be exceeded – warming the sample is part of the process…
  [P] Ability to change fields rapidly during testing session to maximize test time – sufficient – to ramp the magnet with 20-30 A/S (max ~ 150 A/s)
  [D- VD] Fast ramp (1-2 T/s) for pulsed external field testing of joint and
  conductor stability (e.g. for Central Solenoids) is essential but not achievable in this facility - external applied field, not practical to be implemented in the current magnet design, this means an additional specialized magnet to be built and maybe with new high voltage power supplies.

## Slide 6: Recommendations: General facility capabilities
Quench testing capabilities(High priority)
  [M-D] DC power supply capable of ~50 kA -> full evolution of quench - can we use the 30 kA? 50 kA to have? SULTAN have 15kA. This is needed to study current dependence during the quench. 50 kA leads?
  [M-need more investigation] Different quench methods (sample warm-up, AC loss induction, surface heaters)
  [E-M] Quench instrumentation (next slide); helium instrumentation
[D- VD] AC loss testing capability (High Priority) - not practical to do so
  Frequencies for fusion are ~1 HZ but other fields may benefit from up to 200 Hz - +/- 0.4 T not AC loss, not give you full penetration in the stack , 20V 0.15H 150 A/s
  Implement fast-ramp power supply for cable current? Possible to implement oscillating external B-field?

## Slide 7: Recommendations: General facility capabilities
Users will want flexibility to perform Ic (e.g. I vs. V) and Tc (e.g. T vs. V) characterizations
  [M] Excellent control of helium needed (including heaters on inlets): T stability for Ic measurements, at 4 K ~ 10 bar, Tevatron pre-cooler, 2.5 g per second, 8g per leg
[E-M] Flexibility for user-designed cable assemblies
  Clever design of assembly can multiply facility utility (e.g. VIPER cable test at SULTAN)
  External magnet and test well must be sufficiently large and flexible to accommodate - Same as EDIPO
[P] Option to facility-provided sample assemblies
  Not all users may have resources/expertise to design and build assemblies
  FNAL could consider providing “standard assemblies” to users to decrease barrier-to-test - the envision is to have set of assemblies

## Slide 8: Recommendations: Instrumentation and data acquisition
[E-M] Provide a standard “base” instrumentation layout for all tests
  [E] User samples: Voltage taps (high-field, joints), Cernoxes, etc.
  [E] Facility: Voltages (e.g. sample-facility joints), Cryogenics (p,m-dot, etc), B-field, etc.
Flexibility to accommodate user-requested/supplied instrumentation
  [E-M] - depends on the number] Additional voltage taps and Cernoxes specific to user sample
  [M] Surface heaters (with larger gauge wires and in-house power supplies)
  [E] Strain gauges (w/ 3 or 5 wire cable runs), acoustic sensors, fiber optics, hall probes, etc.
  [E-M] Sample cryogenic measurements (pressure, mass flow, etc.)
  [P] Ensure clear instrumentation wire/cable routing from sample to 300 K feedthroughs
  [P- E] Ensure standard feedthroughs (e.g. D-sub) for all facility and user instrumentation
  [P- E] Ensure sufficient spare feedthroughs/flanges beyond “base” instrumentation
  [P- E] Ensure sufficient spare low- (~0.1Hz) and high-rate channels (~100 Hz) in DAQ
  [E-M] Ensure users can bring their own I&C systems for specialized testing needs
  [E-M] Ensure user I&C can tie into facility I&C (esp. timing/synchronization signals)
INTERNAL

## Slide 9: Recommendations: User support and resources – too early to be discussed
Maintain a comprehensive, up-to-date facility manual with all necessary specifications (geometry, interfaces, operational parameters, instrumentation, etc.) for facility users
Generate a post-test summary report (including sample and facility data) and basic set of analyses on sample performance from FNAL to users
Ensure it is straightforward to obtain FNAL-User agreements. Protection of user data and IP-sensitive science and engineering aspects likely critical for privately-funded users.
The following should be provided to maximize user experience and test benefit:
  A clear costing structure for testing + additional labor from FNAL at user’s request (and ensure it is
  competitive with SPC EDIPO-II/SULTAN!)
  A clear scheduling system based on samples delivered at some threshold of readiness
  Detailed shipping/receiving instructions (esp. for international users)
  Access to expert FNAL technicians for sample preparation and installation
  A secure workshop area where users can prepare their own sample (w/ FNAL technicians if needed)
  Dedicated area for users in facility control room (observe/guide test program, analyze data in real time)

## Slide 10: Summary
We got the requests (“wish list”) from the potential users.
Not everything can be done - users asked for test facility that can do everything
We concentrating on:
  the supercritical helium cooling with programmable mass flow rates
  progressing on the design of the transformer
  design of the sample holder.
Other request are in process –
  like raise samples and place joints in-field
  Using the 30kA existing PS to power the samples directly
None of these were initially costed especially the supercritical helium cooling.