## Slide 1: Decay data measurements using MMCs at CEA-LNHB
  Matias Rodrigues
  Martin Loidl
  Arshjot Kaur (PhD student)
  Mostafa Zahir (PhD student)
MiND Workshop, June 28, 2023

## Slide 2: Outline
Introduction
  Presentation of the CEA-LNHB
  Short history of the MMCs at LNHB
Decay data by decay energy spectrometry
  Beta shape of 63Ni and 241Pu
  BR and end-point energy of 151Sm
  Electron capture probabilities of 55Fe
Absolute L X-ray emission intensities of actinides
  The MMC SMX3
  L X-rays from the decay of 238Pu, 244Cm 233Pa, 237Np
  L X-rays from the decay of 241Am

## Slide 3: Introduction
1

## Slide 4: The Laboratoire National Henri Becquerel Activity Metrology
Conventional methods for activity standardization
4πβ-γ coincidence
a-particle counting by defined solid angle
Liquid scintillation (TDCR method)
Triple internal gas counting
And others…
Decay data measurements and evaluations
Beta and gamma-ray spectrometry
Monochromatic sources of photons
Code Beta shape
DDEP Decay Data Evaluation Project (Data recommendation dissemination of data http://www.lnhb.fr/donnees-nucleaires/module-lara/ )
Metallic magnetic calorimeters
  Radioactive source preparation
  Authorizations to prepare and to measure many radionuclides with large activity,
  Liquid, solid or gaseous radioactive source
  Radiochemical separation
  Source activities with low uncertainties

## Slide 5
Puce SQUID
B
SQUID
Magnetic sensor
V
Absorber
Amplification
MMC at LNHB: started in 2004 with the direct coupling
SQUID
Absorber
SQUID chip
Beta spectrum of 241Pu
Electroprecipited 241Pu source (8 Bq) between two 12 µm thick gold foils
Energy calibration: external 55Fe source
Data recorded continuously 2.6 days @ 16 mK
Coil for the field B
1 mm
X-ray
55Fe
241Pu
Calculation

## Slide 6: MMC at LNHB : 2008, Coupling with meander shape pick-up coil
Thermal bath Tbath ~ 30 mK
SQUID
V
Input coil
Gold absorber
πx0.3x0.3x1 mm3
Sensor
I0
senseur
23
A. Burck et al., J Low Temp Phys (2008) 151: 337–344
DEFWHM = 52 eV @ 30 keV
K X-ray spectrum of 133Ba
MMC chip produced by the KIP Heidelberg
sensor
Thermal
link
SQUID
1 mm

## Slide 7: MMC chips produced in collaboration with KIP Heidelberg 2010-2021 and KIT Karlsruhe 2021-now
Different MMC sizes produced to address different projects (European project of metrology MetroBeta, MetroMMC, PrimA-LTD)
25
- 2012 MetroMMC
(249 μm)²
7.57 pJ/K
(335 mm)²
28 pJ/K
(538 μm)²
(1427 μm)²
(2663 μm)²
17.9 pJ/K
308 pJ/K
1.17 nJ/K
sensors
Thermal links
7.2 mm
For a given radionuclide and a given spectrometry, absorber sized is to achieve the required efficiency.
  The sensor MMC size chosen such as the sensor matches the absorber heat capacity
  The source+absorber prepared at LNHB and attached to the sensor by gluing

## Slide 8: Decay data by decay energy spectrometry
2

## Slide 9: Beta spectrum shapes of 241Pu and 63Ni
The measured spectra have:
100% detection efficiency
Energy threshold
  200 eV for 63Ni, 300 eV for 241Pu
FWHM energy resolution
  51 eV for 63Ni, 29 eV for 241Pu
shown evidence of the exchange effect
validated the theoretical code BetaShape
241Pu and 63Ni are pure b emitters to the ground state (GS)
Activity measurements only is achieved by LSC
The knowledge of the b shape is needed
Electro-deposited sources
Energy calibration source
M. Loidl, et al. ARI Vol. 84, P302, 2014
https://doi.org/10.1016/j.apradiso.2013.11.024
M. Loidl et al. ARI Vol. 68 P1454, 2010
https://doi.org/10.1016/j.apradiso.2009.11.054
Photon
X. Mougeot, EPJ Web Conf. Vol. 146, 2017
https://doi.org/10.1051/epjconf/201714612015

## Slide 10: Impact on the activity measurement by LSC
K. Kossert, X. Mougeot, ARI Vol 101, 2015, P40-43
https://doi.org/10.1016/j.apradiso.2015.03.017
Counting efficiency in liquid scintillation depends on the beta spectrum for the 2 LSC methods:
  CIEMAST/NIST (CN) (2 PMTs)
  TDCR (triple to double coincidence ratio) (3 PMTs)
Better agreement between the 2 LSC methods taking into account the exchange effect
Birks’s expression
CN
TDCR
W exchange
W/O exchange

## Slide 11: Beta spectrum of 151Sm branch probabilities and end-point
K. Kossert, ARI, Vol. 185, 110237, 2022
https://doi.org/10.1016/j.apradiso.2022.110237
Curie plot
MiND 2023, June 27-28 2023 | Matias Rodrigues
Beta to excited state
Beta to ground state
Measured Pβ,ground = 99.31 (11)% and Pβ,excited = 0.69 (11)%
Recommended value DDEP: 99.07 (4)% and 0.93 (4)%
Measured Emax = 76.430 (68) keV.
Recommended value AME2020: 76.5 (5) keV
ROIβ,ground
ROIβ,excited
Fit

## Slide 12: Electron capture probabilities (EC) of 55Fe
Like pure beta emitters to the GS, EC radionuclides to the GS can only be measured in LSC
EC probabilities is required.
Source prepared by electroplating on Au foil
Source foil with ~ 10 Bq of 55Fe between two Au foils
Absorber dimensions 600 µm × 600 µm × 24 µm
M. Loidl, et al. ARI, Vol. 134, P395, 2018
https://doi.org/10.1016/j.apradiso.2017.10.042
EC probability | Calculated Value | Pengra et al. 1972 | Loidl et al. 2018 PN = 0.0014 (sstat.)
PK | 0.8853 (16) | 0.881 (4) | 0.8833 (26)
PL | 0.0983 (13) | 0.103 (4) | 0.1001 (22)
PM | 0.0157 (6) | 0.0161 (8) | 0.01515 (38)
K Captures
(6,539 keV)
L Captures
(770 eV)
M Captures
(84 eV)
1 week data taking i.e. 4 × 106 counts
Absorber with 55Fe

## Slide 13: Decay data by photon spectrometry
3

## Slide 14: “Absolute” photon emission intensities (PEIs)
Absolute PEIs: = number of photons at E per 100 decays
Essential decay parameter for quantitative analysis by photon spectrometry
Absolute PEIs are challenging to measure accurately with standard deviation < 1%...
is the product of where
and are difficult to calibrate accurately
Spectrometer
Source
Photon
Energy (keV)
Counts
5
15
25
35
45
55
65
1
10
100
1000
10000
: counts in the Full Energy Peak (FEP)
: FEP detection efficiency
: source activity
: live time
: geometrical factor between source-collimator-absorber
: intrinsic detection efficiency
MiND 2023, June 27-28 2023 | Matias Rodrigues

## Slide 15: SMX3: A dedicated MMC for L X-ray spectrometry of actinides
KIP Heidelberg
V
Thermal bath T0 ~ 10 mK
SQUID
V
Input
coil
4 absorbers of 1 mm2
  50 mm of Au + 17 mm of Ag thick
Intrinsic efficiency > 99% between 10-25 keV
10 – 20 s-1 (td ≈ 4 ms)
Energy resolution FWHM of 22 - 40 eV
Au-Ag absorber
1 mm2
Sensor
AuAg
absorber
SQUID
chip
Thermal
link
MMC with 4 pixels
2 mm
I0
sensor

## Slide 16: Intrinsic Detection efficiency of AuAg absorbers
Source
collimator
X
Ag
absorber
Source
collimator
X
Au-Ag
absorber
Source
collimator
X
Au
absorber
L X-rays for Z > 75
Quasi-constant intrinsic efficiency below 25 keV,
Minimize the efficiency correction
28

## Slide 17: Full energy peak detection efficiency calibration using 241Am and MC simulations
Efficiency calibration
241Am spectrum to establish experimental data of
Extendable dead applied to MMCs to determine
Definition of a meta-geometry by and for Monte Carlo simulations
Efficiency uncertainty
< 10 keV
  ~ 4% uncertainty at 2 keV given by Be window
10 keV < E < 25 keV
  0.7% uncertainty given by
> 25 keV
  1.2% to 2.4% uncertainty given by
10 mm
2 mm
collimator
Absorbers
34

## Slide 18: L X-ray emission intensities of actinides
Many actinides decay by a-emissions and have intense L X-ray PEIs between 10-25 keV
Nuclide | LX-ray intensity / 100 dis. | relative unc. (%) | Method | Number of measurements |  | calc. vs. exp.
Pu-239  U-235+a | 4.66 | 1.1 | meas. | 4 |  | disagree
Pu-240  U-236+a | 10.34 | 1.5 | meas. | 6 |  | agree
Am-241  Np-237+a | 37.66 | 0.5 | meas. | 9 |  | disagree
Cm-242  Pu-238+a | 9.92 | 2.3 | calc. | 2 |  | disagree
Cm-244  Pu-240+a | 8.92 | 2.6 | calc. | 1 |  | agree
U-235  Th-232+a | 40.0 | 55.0 | calc. | - |  | -
241Am L X-ray spectrum
La
Lb
Lg
Li
26
18
238Pu
Chart: Deviation from recommended value (%)
| Halley 1964 | Byrne 1968 | Salgueiro 1968 | Swinth 1968 | Swinth 1971 | Valsiliki 1976 | Bemis 1977 | Bortels 1984 | Dryak 1984 | Johnston 1995 | Schönfeld 2000 | This work
 | -0.0028222 | 0.20696 | -0.13452 | 0.053622 | 0.072437 | 0.33396 | 0.070555 | 0.086548 | -0.00094073 | 0.0047037 | -0.0075259 | 0.0084666

## Slide 19: Data and spectrum processing
Counts
16100
17100
18100
19100
1
10
100
1000
10000
Energy (eV)
20500
21000
21500
22000
22500
23000
1
10
100
1000
10000
11200
12200
13200
14200
1
10
100
1000
10000
Li
Lt
Ls
La1
La2
Lh
Lb6
Lb2,15
Lb5
Lb1
Lb3
Lb9
Lg5
Lg1
Lg6
Au Lb1
W Lg1
4810
5010
5210
10
100
L1-L3
Each spectrum measured during 10-15 days
Data continuously recorded with a 16 bits resolution DAQ @ 250-500 kHz
Data analyzed offline
  Pulse triggering
  Live time determination
  Pulse energy estimation
  Pile-up rejection
  Temperature drift correction
  Non-linearity correction
  Energy resolution equalization
  Spectrum co-adding
Counting statistics of few 106 X-ray counts
  FWHM at best of 22 eV and 32 eV between 0 and 166 keV)
Spectrum processing
  Diagram (solid lines) and satellites (dashed lines)
R. Mariam et al. Spectrochim. Acta B Vol. 187, 2022, 106331
https://doi.org/10.1016/j.sab.2021.106331
19
244Cm

## Slide 20
 |  | 238Pu(a) → 234U |  |  | 244Cm(a) → 240Pu |  |  | 233Pa(b-) → 233U | 
Group |  | This work | Johnston [34] |  | This work | Johnston [34] |  | This work | Calculated
Li |  | 0.2418 (29) | 0.231 (3) |  | 0.2306 (35) | 0.214 (3) |  | 1.075 (19) | 1.05 (4)
La |  | 3.816 (43) | 3.81 (3) |  | 3.49 (5) | 3.38 (3) |  | 15.69 (20) | 16.9 (6)
Lh |  | 0.1284 (16) | 0.126 (2) |  | 0.1002 (22) | 0.102 (2) |  | 0.235 (19) | 0.272 (16)
Lb |  | 5.23 (6) | 5.18 (4) |  | 4.22 (6) | 4.08 (3) |  | 16.89 (25) | 18.1 (6)
Lg |  | 1.291 (14) | 1.29 (1) |  | 1.023 (15) | 0.991 (8) |  | 3.97 (6) | 4.23 (14)
Actinide |  | This work (%) |  | Evaluated or other measured values (%) |  | 
238Pu |  | 10.72 (11) |  | 10.62 (32) | [35] | 
 |  |  |  | 10.63 (8) | [34] | 
244Cm |  | 9.08 (16) |  | 8.77 (6) | [34] | 
 |  |  |  | 9.44 (10) | [36] | 
237Np |  | 56.5 (7) |  | 59.7 (32) | [9] | 
 |  |  |  | 54.8 (21) | [40] | 
Agreement k = 1
Agreement k = 2
In disagreement
Comparison with PEIs in the literature
Total L X-ray emission intensity
Siegbahn group L X-ray emission intensities
Chart: Deviation (%)
| Halley 1964 | Byrne 1968 | Salgueiro 1968 | Swinth 1968 | Swinth 1971 | Valsiliki 1976 | Bemis 1977 | Bortels 1984 | Dryak 1984 | Johnston 1995 | Schönfeld 2000 | This work
 | -0.0028222 | 0.20696 | -0.13452 | 0.053622 | 0.072437 | 0.33396 | 0.070555 | 0.086548 | -0.00094073 | 0.0047037 | -0.0075259 | 0.0084666
238Pu(a) → 234U
See references in https://doi.org/10.1016/j.sab.2021.106331

## Slide 21
Line |  |  |  | 238Pu(a) → 234U |  |  |  |  |  |  |  |  |  | 244Cm(a) → 240Pu |  |  |  |  |  |  |  | 
Siegbahn | IUPAC |  |  | Energy |  | X-ray emission intensity |  |  |  | Rel. Unc. (%) |  |  |  | Energy |  | X-ray emission intensity |  |  |  | Rel. Unc. (%) |  | 
 |  |  |  | (eV) |  | per 100 L X-rays |  | per 100 decays |  | u1 | u2 | u3 |  | (eV) |  | per 100 L X-rays |  | per 100 decays |  | u1 | u2 | u3
- |  | L1-L3 |  | 4589.2 |  | 0.0214 (44) |  | 0.00229 (47) |  | 0.72 | 20 | 1.2 |  | 5054.63 |  | 0.0267 (13) |  | 0.00242 (13) |  | 0.37 | 5.0 | 1.6
Li |  | L3-M1 |  | 11618.4 |  | 2.257 (11) |  | 0.2418 (30) |  | 0.72 | 0.49 | 0.86 |  | 12124.4 |  | 2.540 (13) |  | 0.2306 (36) |  | 0.37 | 0.49 | 1.4
Ls |  | L3-M2 |  | 11982.0 |  | 0.0227 (15) |  | 0.00243 (16) |  | 0.72 | 6.6 | 0.86 |  | 12503 |  | 0.0269 (20) |  | 0.00244 (18) |  | 0.37 | 7.4 | 1.4
Lt |  | L3-M3 |  | 12864.7 |  | 0.0286 (17) |  | 0.00307 (18) |  | 0.72 | 5.8 | 0.86 |  | 13485.4 |  | 0.0315 (16) |  | 0.00286 (15) |  | 0.37 | 5.2 | 1.4
La2 |  | L3-M4 |  | 13439.8 |  | 3.604 (7) |  | 0.3862 (44) |  | 0.72 | 0.20 | 0.86 |  | 14074.5 |  | 3.806 (11) |  | 0.346 (5) |  | 0.37 | 0.29 | 1.4
La1 |  | L3-M5 |  | 13614.8 |  | 32.007 (20) |  | 3.429 (39) |  | 0.72 | 0.062 | 0.86 |  | 14237 |  | 34.68 (7) |  | 3.149 (47) |  | 0.37 | 0.21 | 1.4
Lh |  | L2-M1 |  | 15398.8 |  | 1.199 (6) |  | 0.1284 (16) |  | 0.72 | 0.50 | 0.86 |  | 16333.5 |  | 1.103 (18) |  | 0.1002 (22) |  | 0.37 | 1.6 | 1.4
Lb6 |  | L3-N1 |  | 15726.1 |  | 0.593 (9) |  | 0.0636 (12) |  | 0.72 | 1.5 | 0.86 |  | 16497.9 |  | 0.635 (7) |  | 0.0576 (11) |  | 0.37 | 1.2 | 1.4
Lb2,15 |  | L3-N4.5 |  | 16386.8 |  | 7.626 (10) |  | 0.817 (9) |  | 0.721 | 0.129 | 0.858 |  | 17207.5 |  | 8.271 (19) |  | 0.751 (11) |  | 0.37 | 0.23 | 1.4
 |  |  |  | 16539.6 |  |  |  |  |  |  |  |  |  | 17411.8 |  |  |  |  |  |  |  | 
Lb4 |  | L1-M2 |  | 16575.0 |  | 0.2240 (29) |  | 0.02400 (41) |  | 0.72 | 1.3 | 0.86 |  | 17555.2 |  | 0.2506 (47) |  | 0.0228 (5) |  | 0.37 | 1.9 | 1.4
Lb17 |  | L2-M3 |  | 16647.1 |  | 0.0681 (25) |  | 0.00730 (28) |  | 0.72 | 3.7 | 0.86 |  | 17702.8 |  | 0.0531 (13) |  | 0.00482 (14) |  | 0.37 | 2.5 | 1.4
Lu |  | L3-N6,7 |  | 16774.5 |  | 0.257 (11) |  | 0.0276 (12) |  | 0.72 | 4.3 | 0.86 |  | 17629.1 |  | 0.240 (8) |  | 0.0218 (8) |  | 0.37 | 3.5 | 1.4
Lb7 |  | L3-O1 |  | 16974.0 |  |  |  |  |  |  |  |  |  | 17749.4 |  |  |  |  |  |  |  | 
Lb5 |  | L3-O4,5 |  | 17064.0 |  | 1.636 (31) |  | 0.1753 (38) |  | 0.72 | 1.87 | 0.86 |  | 17951.2 |  | 1.725 (14) |  | 0.1566 (26) |  | 0.37 | 0.81 | 1.4
 |  | L3-P1,2,3 |  | 17123.2 |  |  |  |  |  |  |  |  |  | 18072.4 |  |  |  |  |  |  |  | 
Lb1 |  | L2-M4 |  | 17220.2 |  | 38.19 (11) |  | 4.091 (47) |  | 0.72 | 0.28 | 0.86 |  | 18293.3 |  | 35.138 (34) |  | 3.190 (47) |  | 0.37 | 0.10 | 1.4
Lb3 |  | L1-M3 |  | 17455.0 |  | 0.1933 (27) |  | 0.02071 (37) |  | 0.72 | 1.4 | 0.86 |  | 18541 |  | 0.1879 (27) |  | 0.01706 (35) |  | 0.37 | 1.4 | 1.4
Lb10 |  | L1-M4 |  | 18029.3 |  | 0.01298 (31) |  | 0.001390 (36) |  | 0.72 | 2.4 | 0.86 |  | 19137 |  | 0.0045 (6) |  | 0.00041 (6) |  | 0.37 | 14 | 1.4
Lb9 |  | L1-M5 |  | 18204.9 |  | 0.01711 (35) |  | 0.001833 (43) |  | 0.72 | 2.1 | 0.86 |  | 19330.3 |  | 0.0122 (8) |  | 0.00111 (7) |  | 0.37 | 6.5 | 1.4
Lg5 |  | L2-N1 |  | 19506.4 |  | 0.3324 (45) |  | 0.0356 (6) |  | 0.72 | 1.3 | 0.86 |  | 20707.4 |  | 0.3154 (34) |  | 0.0286 (5) |  | 0.37 | 1.1 | 1.4
- |  | L2-N3 |  | 19901.0 |  | 0.0242 (16) |  | 0.00259 (18) |  | 0.72 | 6.7 | 0.86 |  | 21146.6 |  | 0.0242 (11) |  | 0.00220 (10) |  | 0.37 | 4.4 | 1.4
Lg1 |  | L2-N4 |  | 20166.5 |  | 9.386 (8) |  | 1.006 (11) |  | 0.72 | 0.088 | 0.86 |  | 21416.8 |  | 8.683 (24) |  | 0.788 (12) |  | 0.27 | 0.27 | 0.3
Lg2 |  | L1-N2 |  | 20482.0 |  | 0.0615 (10) |  | 0.00659 (13) |  | 0.72 | 1.7 | 0.86 |  | 21720.6 |  | 0.0705 (32) |  | 0.00640 (31) |  | 0.37 | 4.6 | 1.4
Lg'8 |  | L2-N6 |  | 20556.1 |  | 0.0616 (8) |  | 0.00660 (11) |  | 0.72 | 1.3 | 0.86 |  | 21829.3 |  | 0.0607 (21) |  | 0.00551 (21) |  | 0.37 | 3.4 | 1.4
Lg8 |  | L2-O1 |  | 20625.7 |  | 0.0910 (11) |  | 0.00975 (16) |  | 0.72 | 1.2 | 0.86 |  | 21914 |  | 0.0646 (18) |  | 0.00586 (18) |  | 0.37 | 2.8 | 1.4
Lg3 |  | L1-N3 |  | 20712.4 |  | 0.0594 (13) |  | 0.00636 (15) |  | 0.72 | 2.1 | 0.86 |  | 21980 |  | 0.0665 (18) |  | 0.00604 (18) |  | 0.37 | 2.6 | 1.4
Lg6 |  | L2-O4 |  | 20842.0 |  | 1.948 (6) |  | 0.2087 (18) |  | 0.72 | 0.33 | 0.36 |  | 22149.1 |  | 1.911 (9) |  | 0.1735 (27) |  | 0.37 | 0.45 | 1.4
- |  | L2-P1 |  | 20904.0 |  | 0.0117 (17) |  | 0.00126 (18) |  | 0.72 | 15 | 0.86 |  | - |  | - |  | - |  | - | - | -
- |  | L2-P4,5 |  | 20941.7 |  | 0.03312 (50) |  | 0.00355 (7) |  | 0.72 | 1.5 | 0.86 |  | 22260.9 |  | 0.0343 (11) |  | 0.00312 (11) |  | 0.37 | 3.2 | 1.4
Lg'4 |  | L1-O2 |  | 21498.1 |  | 0.01838 (41) |  | 0.001969 (49) |  | 0.72 | 2.2 | 0.85 |  | 22823.3 |  | 0.0141 (7) |  | 0.00128 (7) |  | 0.37 | 5.0 | 1.4
Lg4 |  | L1-O3 |  | 21564.0 |  | 0.01345 (34) |  | 0.001441 (39) |  | 0.72 | 2.5 | 0.85 |  | 22888.8 |  | 0.0162 (8) |  | 0.00147 (7) |  | 0.37 | 4.6 | 1.4
- |  | L1-P2 |  | 21736.6 |  | 0.00561 (27) |  | 0.000601 (30) |  | 0.72 | 4.9 | 0.85 |  | 23082.1 |  | 0.0059 (5) |  | 0.000537 (46) |  | 0.37 | 8.5 | 1.4
Tens of PEIs of individual L X-ray transitions

## Slide 22: Limits of previous measurements
Previous PEI measurements limited by the uncertainty on the FEP detection efficiency
Itself limited by the recommended values of the PEIs of 241Am used for the efficiency calibration
PEIs are ultimately interdependent and correlated to some extent

## Slide 23: Absolute L X-ray PEIs of 241Am independently of other PEIs
: rate of photons in the FEP
: rate of a-particle in the FEP
: FEP efficiency for photons
: efficiency for a particle
: intrinsic efficiency
: geometrical efficiency
Conditions:
  100% decay by a-particle emission
However measuring a-particles and photons in the same spectrum is not possible:
  MMC sensitivity for a-particles (few MeV) and X-rays (10s of keV) must be different
  Ten times as many emitted particles per decay as there are X-rays
     Many pile-up or reduced counting statistics, electrons/X-ray interferences
Spectrometer
241Am source
Photon
a

## Slide 24: Measurements of two 241Am sources with different activities

1st 241Am source
collimator
absorber
MMC SMX3
10 mm
1st measurement: a spectrum
241Am source of 1.8 kBq
Lower MMC sensitivity
No Be window
FWHM resolution of 3.3 keV
Energy / keV
Counts per 0.4 keV
5000
5100
5200
5300
5400
5500
1
10
100
1000
10000
a-particle spectrum

## Slide 25: Measurements of two 241Am sources with different activities
Be window
2nd source
collimator
MMC SMX3
10 mm
Energy / keV
Counts per 0.4 keV
5000
5100
5200
5300
5400
5500
1
10
100
1000
10000
1st measurement: a spectrum
241Am source of 1.8 kBq
Lower MMC sensitivity
No Be window
FWHM resolution of 3.3 keV
2nd measurement: X-ray spectrum
241Am source of 32 kBq
High MMC sensitivity
Be window to stop the a-particles
spectrum FWHM resolution of 28 eV
Energy (eV)
241Am a-particle spectrum
L X-ray spectrum of 241Am(a) → 237Np



## Slide 26: Additional corrections due to the measurement of 2 sources
and from energy spectra
ratio between source activities determined by conventional a-particles spectrometry
correction factor for the inhomogeneity
of the surface source activitiy determined by
radioactive source imager.
and intrinsic efficiencies ~ 1, determined by Monte Carlo simulations.
transmission through Be window, calculated.

## Slide 27: Results of L X-ray PEIs from 241Am(a) → 237Np
Total L X-ray and Siegbahn group (La, Lb, Lg…) PEIs
  Good agreement with all the previous published data
  Relative uncertainty (0.32%) 2 times lower than the most precise measurement

## Slide 28: Results of L X-ray PEIs from 241Am(a) → 237Np
33 PEIs of individual X-ray transitions are provided
X-ray transition |  | Energy (eV) | PEI per 100 decays | Relative uncertainties |  |  | Siegbahn group | PEI per 100 decays | Relative uncertainties |  | 
UIPAC | Siegbahn |  |  | Total | Fitting procedure | Counting statistics |  |  | Total | Fitting procedure | Counting statistics
L1-L3 | - | 4820 | 0.2289 (15) | 0.65% | 0.41% | 0.42% |  |  |  |  | 
L3-M1 | Li | 11873 | 0.8989 (38) | 0.42% | 0.13% | 0.21% |  |  |  |  | 
L3-M2 | Lt | 12250 | 0.01024 (26) | 2.57% | 0.64% | 1.96% |  |  |  |  | 
L3-M3 | Ls | 13179 | 0.0101 (8) | 7.59% | 7.33% | 1.98% |  |  |  |  | 
L3-M4 | La1 | 13762 | 1.2581 (49) | 0.39% | 0.19% | 0.18% | La | 13.046 (41) | 0.31% | 0.084% | 0.069%
L3-M5 | La2 | 13944 | 11.788 (38) | 0.32% | 0.13% | 0.06% |  |  |  |  | 
L2-M1 | Lh | 15862 | 0.3745 (23) | 0.60% | 0.29% | 0.33% |  |  |  |  | 
L3-N1 | - | 16113 | 0.2389 (14) | 0.59% | 0.32% | 0.41% | Lb | 18.68 (6) | 0.31% | 0.064% | 0.058%
L3-N4,5 | Lb2,15 | 16796 - 16958 | 2.912 (14) | 0.48% | 0.36% | 0.12% |  |  |  |  | 
L1-M2 | Lb4 | 17061 | 1.602 (5) | 0.33% | 0.04% | 0.16% |  |  |  |  | 
L3-N6,7 | Lb7' | 17207 - 17222 | 0.01636 (49) | 2.99% | 2.54% | 1.55% |  |  |  |  | 
L3-O1,2,3 | Lb7 | 17275 - 17357 | 0.0818 (32) | 3.95% | 3.88% | 0.69% |  |  |  |  | 
L3-O4,5 - L3-P1,4,5 | Lb5 | 17501 - 17630 | 0.638 (9) | 1.45% | 1.39% | 0.25% |  |  |  |  | 
L2-M4 | Lb1 | 17750 | 11.665 (39) | 0.34% | 0.16% | 0.06% |  |  |  |  | 
L2-M5 | - | 17938 | 0.00976 (31) | 3.21% | 2.49% | 2.01% |  |  |  |  | 
L1-M3 | Lb3 | 17992 | 1.3242 (49) | 0.37% | 0.14% | 0.17% |  |  |  |  | 
L1-M4 | Lb10 | 18577 | 0.0756 (6) | 0.82% | 0.25% | 0.72% |  |  |  |  | 
L1-M5 | Lb9 | 18762 | 0.1150 (8) | 0.69% | 0.21% | 0.59% |  |  |  |  | 
L2-N1 | Lg5 | 20101 | 0.1037 (10) | 0.93% | 0.63% | 0.62% | Lg | 4.883 (20) | 0.41% | 0.26% | 0.11%
L2-N3 | - | 20516 | 0.0056 (9) | 16.41% | 16.19% | 2.65% |  |  |  |  | 
L2-N4 | Lg1 | 20785 | 2.916 (12) | 0.41% | 0.20% | 0.12% |  |  |  |  | 
L1-N2 | Lg2 | 21098 | 0.4330 (20) | 0.46% | 0.19% | 0.30% |  |  |  |  | 
L2-N6 | Lg8' | 21186 | 0.01811 (31) | 1.72% | 0.83% | 1.48% |  |  |  |  | 
L2-O1 | Lg8 | 21260 | 0.02841 (48) | 1.69% | 1.17% | 1.18% |  |  |  |  | 
L1-N3 | Lg3 | 21341 | 0.4363 (20) | 0.46% | 0.20% | 0.30% |  |  |  |  | 
L2-O4 | Lg6 | 21489 | 0.6260 (30) | 0.48% | 0.28% | 0.25% |  |  |  |  | 
L2-P1 | - | 21555 | 0.00386 (16) | 4.02% | 2.41% | 3.20% |  |  |  |  | 
L2-P4 | - | 21595 | 0.02246 (31) | 1.39% | 0.28% | 1.33% |  |  |  |  | 
L1-N5 | - | 21656 | 0.02041 (30) | 1.46% | 0.32% | 1.39% |  |  |  |  | 
L1-O2 | Lg4' | 22155 | 0.1139 (8) | 0.67% | 0.15% | 0.59% |  |  |  |  | 
L1-O3 | Lg4 | 22216 | 0.1057 (7) | 0.70% | 0.15% | 0.61% |  |  |  |  | 
L1-O4,5 |  | 22319 | 0.00639 (16) | 2.58% | 0.61% | 2.49% |  |  |  |  | 
L1-P2,3 |  | 22404 | 0.04361 (46) | 1.05% | 0.34% | 0.95% |  |  |  |  | 

## Slide 29: Conclusions and perspectives
MMCs are suitable for decay data measurements by DES (< 200 keV)
  Beta spectrum shape
  End point energies
  Beta branch probabilities
  EC probabilities
   We are currently developing multiple MMC channels to achieve high statistics (108 counts) for beta spectrum of 129I
MMCs are suitable for precise and detailed absolute PEI determinations (<100 keV)
  L X-ray PEIs for 238Pu, 244Cm, 233Pa and 237Np with relative uncertainties of ~ 0,8%
  L X-ray PEIs for 241Am without efficiency calibration based on other PEIs with relative uncertainties of ~0.3%
  g-ray and X-ray PEIs in the range 25 keV-100 keV with relative uncertainties of ~ 0,8%-1%
   We are currently developing an MMC array for absolute PEIs of photons < 10 keV
Measuring these decay data at higher energies is challenging:
  due to loss of Bremsstrahlung photons for electrons
  due to loss of efficiency for photons
MMCs are useful tool to provide accurate decay data in metrology of ionizing radiation

## Slide 30: Thank you for your attention
Matias Rodrigues