## Slide 1: Overview of 1K TEM Project
Lucas Brouwer, Tengming Shen, Ryan Norris, Aurelio Hafalia, Ross Schlueter, Li Wang, Jim Ciston, Peter Ercius, Qing Ji, Marian Mankos, Colin Ophus, Alexander Stibor, Andreas Schmid, Andrew Minor, and Peter Denes
Cross area project: mix of ATAP, Engineering, Molecular Foundry (Energy Sciences and Physical Sciences)
+ Industry (ElectronOptica)

## Slide 2
https://www.wikipedia.org/
Demonstrator for 1K TEM

## Slide 3
https://www.wikipedia.org/
Objective Lens
iron dominated (~1.5 T)
precise pole piece shape determines field
****Pass around example pole pieces
Thin sample disk

## Slide 4: Why do we want to make an all cold microscope?
stability is a key challenge for modern microscopes, amplified by recent advances in aberration correction and the desire for experiments with cryogenically cooled samples
issue | solution explored by 1K-TEM project at LBNL
mechanical drift | cool entire microscope to LHe temp
mag. field drift | ultra-stable, persistent current SC magnets
 | superconducting magnetic shielding
http://www.science.uwaterloo.ca/~gsciaini/research.html
Cryogenic sample holder to place cold sample into a warm microscope

## Slide 5
https://link.springer.com/article/10.1557/mrs.2019.288

## Slide 6: 1K TEM project history
LDRD (1-2 years)
  design study (focused on SC magnet / electron source)
Initial Basic Energy Sciences (BES) funding (~3 years)
  demonstrate SC lens and electron source technology
  build and operate SC lens in persistent current mode and characterize stability
  measuring stability of field with probes
New BES funding (3 years starting now)
  combine electron source, SC lens, and detector into mini-microscope
  measuring stability of beam with detector

## Slide 7: Persistent current mode with a superconducting switch (common for MRI)
SC lens
Cold
+
-
Charging Mode
Rpcs ~1-2 Ohm
Cold
+
-
Persistent Mode
Rpcs ~0
SC lens
SC joint (~pico-Ohm)
SC joint

## Slide 8: SC joints – typically 10-12 Ohm for Nb-Ti

## Slide 9: Nb-Ti persistent current switch
Cowound heater wire
Nb-Ti (Ni-Cu matrix)
Bobbin / caps
Final switch

## Slide 10
Persistent: PCS heater off
Rpcs ~0
SC lens
joint resistance (nΩ -pΩ)
index loss
L/R time constant vs. inductance (assume 1 pΩ resistance)

## Slide 11
We can’t achieve the stability we want need with only persistent current operation, we need active stabilization too

## Slide 12: 1.95 T, prototype Lens

## Slide 13: Magnet assembly with PCS, joints and tuner

## Slide 14: Persistent current operation at 2 T routinely established without quench
Header with magnet assembly
LHe bath cryostat
1.82 T at hall probe
Persistent current operation
Measured field vs. current is in agreement with Opera3D (shows saturation effects)

## Slide 15
We can’t achieve the stability we want need with only persistent current operation, we need active stabilization too

## Slide 16: Movement of the cryogenic stage changes the inductance of the “tuner” magnet which shifts flux to or from the main lens (active tuner to stabilize the current)
iron
series coil
gap change
attocube
SC joint
SC joint
objective
lens
tuner
gap

## Slide 17
https://www.capacitec.com/
Capacitive gap sensors independently measure stage position in LHe
4 K calibration allows for moving to a desired gap change during active stabilization programs
Cap sensor targets
Cap probe mounts
ANPx101 with 5 mm range

## Slide 18: Experimental demonstration: stabilizing to the noise level of the Hall probes (a few ppm at 1.95 T)
Brouwer et al, “Stabilization and control of persistent current magnets using variable inductance”, Supercond. Sci. Technol. 2022, https://doi.org/10.1088/1361-6668/ac549b

## Slide 19: Summary of first phase of 1K TEM
demonstrated we can build a Nb-Ti lens and operate in persistent current mode
new technique for stabilization of persistent current magnets, demonstrated it works up to the precision with which we can measure (in theory it can do much better)
not completed
  magnetic shielding / ultimate stability measurements with SQUIDs etc.
  cryostat stabilization

## Slide 20: Next Steps (3 years starting now)
add magnetic shielding and cryogenic stabilization to our existing experiment
build and test beam-based demonstrator, combine lens with electron source, high vacuum, detector, shielding etc. (all in liquid helium bath)