## Slide 1: Quench analysis for the FRIB Nb-Ti superconducting ECR magnet
Ye Yang1), Tengming Shen1), Paolo Ferracin1), Diego Arbelaez1), Soren Prestemon1), Xiaoji Du2), David Greene2), 3), Danlu Zhang2), Yoonhyuck Choi2), Ting Xu2)
1) Lawrence Berkeley National Laboratory (LBNL), Berkeley, California, USA
2) Facility for Rare Isotope Beams (FRIB), East Lansing, Michigan, USA
3) Michigan State University, East Lansing, Michigan, USA
ECR Workshop, 11/14/2024

## Slide 2: Contents
01
Introduction
02
Quench protection for FRIB-I & Testing result
03
Quench simulation & Benchmarking study
04
Eddy current induced mechanical issue in cryostat
05
Summary

## Slide 3: Introduction
Superconducting ECR magnet for the FRIB consists of 9 coils and iron pole
  6 individual sextupole coils, radial field
  3 solenoid coils, axial field
ECR magnetic field needs to be tunable to generate the desirable charge status and beam intensity
  4 power supplies
  3 independent circuits
Challenge for the quench protection:
  How to protect the multipole coil system?
    Difficult to predict the quench propagation
  Impact of the mutual inductance
  Eddy current induced mechanical issues
Extraction
Middle
Injection
Sextupole
Iron pole
Coil 1
Coil 2
Mandrel
LHe Vessel
Vacuum Vessel
Thermal shield

## Slide 4: Quench protection circuit for FRIB-I
D. Arbelaez, et al, Test results for a superconducting 28-GHz ion source magnet for FRIB, IEEE Trans. Appl. Supercond., 29 (5), 4100605, 2019
Total stored energy, 660kJ
Protected by the internal dump resistor
No switch between the internal and PS
Current also transport to the PS dump resistor during the quench
PS dump resistor
diode
Internal resistor

## Slide 5: Measured current decays and voltages
Time constants derived from the measurement
  Sextupole, 1.8s
  Solenoid, 4.0s
The current decays are all dominated by the quench propagation
Quench in each single sextupole increase the current in the others due to the mutual inductance
The magnetic coupling between the solenoid is not so strong
Reverse diodes are switched on for
  coil4,
  coil3,
  coil7,
  minimum
1.8s
4.0s
Direction reversed

## Slide 6: Quench simulation
Time loop
Electromagnetic simulation
Thermal simulation
Circuit simulation
End of time loop
User defined function
Extract , , and at each element
Calculate
Calculate effective resistivity
Evaluate coupling loss power density
Saving the power density as element heat
OPERA QUENCH
Quench control
Input heat in duration of
If coil voltage >
Quench event is detected
Open S2; Close S1 after
End if
Dump resistor
Power Supply
PYTHON

## Slide 7: 3D model
Solenoid mandrel
(A6061)
Banding (A6061)
Insulation (G10)
Injection
Middle
Extraction
Pad (SUS316)
Iron pole (SS1010)
Insulation (Kapton)
Sextupole 6
Sextupole 5
Sextupole 4
Sextupole 3
Film heater
cryostat
cryostat
Quench is triggered by a film heater located at the coil end
Sextupole cold mass is not thermal connected to the solenoid
Use the symmetric model to reduce the computation cost
Item | Unit | Value | 
Coil | - | Sextupole | Solenoid
Bare cross section | mm | 1.80 x 0.90 | 1.57 x 0.88
Insulated cross section | mm | 1.90 x 1.00 | 1.65 x 0.96
Cu ratio | - | 3 | 4
Twist pitch | mm | 19 | 19
RRR | - | 160 | 140

## Slide 8: Simulation result with coupling loss
Reasonable current decay, ~4s
Maximum temperature, < 70K
Coil resistance, > 0.5Ω
  Much higher than the internal resistor
Spike of coil voltage due to the time & geometry mesh
Resistive voltage, < 650V

## Slide 9: Comparing the current decays and maximum temperature with the measurement
Predictions of the sextupole coils decay faster than the measurements;
However, the solenoids are matching the measurement perfectly
  This may be attributed to the half model (electrical circuit is not symmetrical)
Since the coupling between the solenoid and the sextupole coils is very low, the solenoid is mainly triggered by the eddy current generated in the mandrel
* Measurements are estimated from MIITs
Coil | Measurement [K] | Simulation [K]
Sextupole | 76 | 58
Injection | 74 | 65
Middle | 52 | 31
Extraction | 72 | 60
Comparison of the maximum temperatures:

## Slide 10: Effect of the coupling loss
Effect of the coupling loss is significant
Reduce the time constant about 1s in both the sextupole and the solenoid
If the coupling loss is not included, the current in sextupole coil increases 18% during the quench due to the mutual inductance

## Slide 11: Comparing the OPERA with the QUENCH code
QUENCH code
OPERA
Compared the OPERA with the QUENCH code
Both results seem consistent only for the single coil
The QUENCH code cannot handle the mutual inductance and eddy current

## Slide 12: Quench propagation
80
70
60
50
40
30
20
10
4.2
Temperature [K]
0.1s
0.2s
0.5s
0.7s
1.5s
Hot spot
0.1s
0.2s
0.5s
0.7s
1.5s
Quench propagates azimuthally from one sextupole coil to another
Final maximum temperature is located at the place where the maximum field is located

## Slide 13: Eddy current generation during the quench
0.1s
0.2s
0.5s
0.7s
1.5s
10
9
8
7
6
3
2
1
0
Current density [A/mm2]
4
5
0.1s
0.2s
0.5s
0.7s
1.5s
Eddy current is locally generated by the quenched coil
Then, the eddy current heats up the mandrel to trigger the quench in the solenoid

## Slide 14: Net force in cold during the quench
Aluminum mandrel
The force generated in the magnet are going to rebalanced during the quench
Net forces change the sign during the quench due to the quench propagation
Needs the further investigation of the coil stress & contact in future
mandrel
mandrel
mandrel
mandrel
Quench
force

## Slide 15: FRIB ECR magnet cryostat
Item |  | Unit | FRIB
Sextupole coil | IR | mm | 100
GAP (sext  LHe vessel) |  | mm | 2.97
LHe vessel | OR | mm | 97.03
 | IR | mm | 95.25
 | thickness | mm | 1.78
GAP (LHe vessel  thermal shield) |  | mm | 5.08
Thermal shield | OR | mm | 90.17
 | IR | mm | 88.9
 | thickness | mm | 1.27
GAP (thermal shield  vacuum vessel) |  | mm | 4.75
Vacuum vessel | OR | mm | 84.15
 | IR | mm | 82.55
 | thickness | mm | 1.6
GAP (vacuum vessel  plasma chamber) |  | mm | 0.55
Plasma chamber | OR | mm | 82
 | IR | mm | 72
 | thickness | mm | 10
Plasma chamber
Vacuum vessel
Thermal shield
LHe vessel
Thermal shield usually made of the thin, low resistivity material
A quench may destroy the thermal shield due to the eddy current, which may potentially increase the heat leak
Assessment of the quench effect for the cryostat is necessary

## Slide 16: Modeling the quench effect
Measurement data
OPERA/ELEKTRA
ANSYS mechanical
Item | Unit | Thermal shield | LHe vessel | Vacuum vessel
Shape | - | Circular | Circular | Circular
Temperature | K | 50 | 4.2 | 300
Material | - | A1100 | SS316 | SS316
Electric resistivity | Ω·m | 0.16e-8 | 53.9e-8 | 77.1e-8
Inner radius | mm | 88.9 | 95.25 | 82.55
Outer radius | mm | 90.17 | 97.03 | 84.15
Thickness | mm | 1.27 | 1.78 | 1.6
Length | mm | 1078.23 | 1078.23 | 1078.23
* No iron yoke
Modeling the quench effect by passing the current decay to OPERA/ ELEKTRA
Then passing the force to ANSYS to simulate the stress
Using the measurement data instead of the simulation result
Vacuum vessel / Thermal shield,
frictional
LHe vessel / Thermal shield,
frictional

## Slide 17: Eddy current density in the thermal shield
Time, 0.2s
Time, 0.4s
Jmax = 58.3A/mm2
Jmax = 73.5A/mm2
Jmax = 63.1A/mm2
Time, 0.6s
Time, 1.6s
Time, 1.2s
Jmax = 44.1A/mm2
Jmax = 35.0A/mm2
0.6s
1.2s
The eddy current generation corresponds to the quench propagation

## Slide 18: Force density in the thermal shield
0.2s
fmax = 188MN/m3
fmax = 334MN/m3
0.4s
0.6s
fmax = 285MN/m3
1.2s
fmax = 170MN/m3
1.6s
fmax = 143MN/m3
0.6s
1.2s
Maximum force density is generated in the thermal during the sextupole quench
Peak net radial force: 8.3kN
Peak net axial force: 1.4kN

## Slide 19: Von Mises stress and deformation at 0.1s
LHe vessel
Thermal shield
Vacuum vessel
197 MPa
192 MPa
73 MPa
10.3 mm
0.7 mm
1.7 mm
LHe vessel
Thermal shield
Vacuum vessel
Deformation
Von Mises stress
Thermal shield touches and transfer the force to the vacuum vessel
Yield strength, 123 MPa < Maximum stress, 197MPa < Ultimate strength, 318 MPa (A1100)

## Slide 20: Conclusion
A quench simulation is performed for the FRIB Nb-Ti ECR magnet with the implementation of the coupling loss effect
Assuming one of the sextupole coils is quench first, the quench of the solenoid will be triggered by the eddy current loss in the mandrel
Thanks to the quench back from the eddy current, the currents of the sextupole and solenoid coils decay in 1s and 4s after the shut-down of the power supply
During the quench, the forces in the magnet are going to rebalance, which may affect the contact and stress in the magnet
The force generated in the thermal shield is a biggest concern, which causes the deformation of 10mm and maximum stress of 197MPa

## Slide 21: Thank You

## Slide 22: Quench protection for VENUS