## Slide 1: Triton Source at Florida State University
Ashton Morelock

## Slide 2: Introduction to Triton Beams
No (other) facility currently exists in the U.S. to produce pure, low energy (3-17 MeV) triton beams.
  (t,p), (t,n), (t,α), (t,He), t+t
There are a significant number of nuclear reactions involving tritium that are crucial for Stockpile Stewardship, Nuclear Fusion, Nuclear Structure, and Nuclear Astrophysics. For many of these, the cross-sections and the distribution of reaction products are either unknown or uncertain.
Tritium was successfully run in the past (UPenn and McMaster) using sputter (solid state) sources.

## Slide 3: Introduction to Triton Beams
= 12.32 years
H → He + + ; Q() = 18.6 keV
  Travels less than 7 mm in air
  Cannot penetrate dead layer of skin
  Cannot penetrate clothing or gloves
Internal exposure
  Inhalation
  Absorption
  Ingestion
3Li
4Li
5Li
6Li
2H
1H
4H
3H
Neutron
3He
4He
5He
Notes: Radioactive isotope of hydrogen
If tritium is present in a material containing hydrogen, the tritium atom will exchange with the hydrogen atom, forming a tritiated version of the material
Internal exposure can mess with DNA and RNA causing genetic consequences
Interlock and ppe are important
1-2 week biological half-life, and can be flushed by drinking a lot of excess fluids
The decay-β will, however induce the emission of Bremsstrahlung X-rays, mostly from the titanium atoms surrounding tritium. Here, the spectrum will show a prominent peak at the Kα-energy of titanium, 4.8 keV, which, is easy to shield by protective equipment.

## Slide 4: Florida State University Tritium Source
Multi-cathode SNICS source
Titanium sublimation pumps
  1st on potential
  2nd before inflector magnet
Lower beam current: ~20 nA
Higher transmission: >98% through tandem
Further limit and negate any tritium contamination.

## Slide 5: MC-SNICS
Notes: Cs res is heated until Cs vapor is formed
The cathode is cooled and the ionizer is heated
Some Cs condenses onto the cooled cathode and some is ionized by the hot surface
+Cs accelerates towards the cathode, sputtering material on impact
Some materials will preferentially sputter neutral or + particles which will pick up electrons as they pass through the condensed Cs produced a negatively charged beam

## Slide 6: Cathodes
9.5 mm x 6.5 mm tantalum Jacket
3 mm x 2 mm titanium insert
5 tritiated cathodes (outside vendor)
Deuterated cathodes (in house)
Up to 20 spots
6.5 mm
9.5 mm
2 mm
3 mm

## Slide 7: Safety Precautions
No handling of gaseous tritium
  If gaseous tritium is detected, precautions are in place to ensure proper storage and disposal
Negative Pressure Enclosure (fume hood) around source
Tritium Monitor (sensitive to 1 μC/m3 )
Interlock System
Titanium Sublimation Pumps
“Unique” equipment for tritium campaigns
  LE buncher
  Slits
  Beam stops
  Targets
Notes: The total activity of 3000 hours of 20 nA tritium beam corresponds to 70 mCi of activity in the beam particles themselves.
This activity will mostly be deposited at the beam stop of the experimental setup and implanted at significant depths in the material, e.g. to 0.7 mm depth in Tantanlum when delivered at the typical 16 MeV energy. From those implantation depths, essentially all Bremsstrahlung X-rays will be absorbed by the Tantalum material

## Slide 8: Vacuum
Differential pumping design
Two Titanium Sublimation Pumps
  TSP1 conductance ratio: 1000/1
  TSP2 conductance ratio: 500/1
Two Ion Pumps (initial pump down)
Turbo pump after TSP2
~0.25” x 1’ gas restriction tube

## Slide 9: Fume Hood
Fume hood designed and built around the source for tritium evacuation
Exhaust from scroll pump is released within fume hood
Tritium monitor samples all outgoing air

## Slide 10: Interlock System
Exception Condition | Interlock Trip |  | 
 | Level 1: HV Off | Level 2: Vacuum Isol. | Level 3: Power Off
Cage door opened | X |  | 
Vacuum above threshold | X | X | 
Coolant flow supply trip | X | X | X
Coolant flow return trip | X | X | X
Smoke detector alarm | X | X | X
Fume hood flow trip | X | X | X
Tritium monitor trip | X | X | X
Power failure | X | X | X
Emergency button | X | X | X

## Slide 11: Tritium Cathode
Inside the current geometry of a 2 mm x 3 mm Titanium pellet, about 20 Ci of tritium can be baked in per pellet
 100 Ci = 5 Cathodes
Bottle was shipped to American Radiolabeled Chemicals (ARC), baked by ARC, backfilled with Argon, then sent back to us

## Slide 12: Commissioning Attempt: 08/27/2024 & 02/07/2025
Both commissioning attempts released too much tritium and were aborted.
It was determined that the tritiated cathodes could not be unloaded from the bottle in air, rather it must be done in an inert gas atmosphere to prevent a large release of tritium.
A glove box has been installed in a dedicated fumehood that is filled with dry nitrogen to keep the humidity under 5%.
A “dry” container was fabricated to contain the cathode and cathode wheel.
Notes: From this commissioning experience, we realized that we would not be able to proceed with decommissioning the source as planned because once in contact with air, all stainless-steel surfaces would release a substantial amount of tritium.

## Slide 13: Decommissioning (short term)
Flow argon through source to remove elemental tritium and cap off Cs reservoir (exhausted into fumehood).
Flow wet air through the source into two sealed, 5-gallon buckets filled with activated alumina to absorb tritiated water.
Remove TSP1 and cathodes and safely perform source maintenance.

## Slide 14: Decommissioning (long term)
Torion Plasma is to provide us with a dual dryer ITS system.
  2x 8 kg drier beds
  1x 2.5 kg Ni bed
The Ni bed will oxidize any elemental tritium, and the drier beds will capture all water vapor.
The drier beds will be regenerated when full and the tritiated water collected. This will be declared as rad waste.
VAC-ITS
Notes: Maybe mention our previous plan of removing the liner, but do to the most recent commissioning attempt we decided to dispose of the entire thing

## Slide 15: Commissioning Attempt: 10/2025 - Ongoing
Cathode was loaded into cathode wheel using dry nitrogen glove box and transported to MC-SNICS.
First extracted deuteron beam to ensure source was operating properly.
~10 nA of tritons on LE cup at beginning of testing (even after optimizing all parameters).
~3nA of tritons on LE cup after around 7 days of running (only day shifts).
Started with ~300 pA on target at SPS and ended with ~2pA on target at SPS.
Ti(t,p)Ti @ E = 16 MeV @ 15°
C(t,α)B @ E = 16 MeV @ 15°
Tritiated
cathode

## Slide 16: Commissioning Attempt: 10/2025 - Ongoing

## Slide 17: Summary
A dedicated MC-SNICS has been installed at John D. Fox Laboratory
  Complete with Infrastructure, vacuum, fume hood, interlocks
Using a single tritium cathode, commissioning was attempted twice
  Too much tritium was released, so both attempts were aborted.
  Procedures have been established to prevent gaseous tritium release and tritium release from cathode bottle.
  Procedures are also being developed to decontaminate the source after commissioning.
Third commissioning attempt is ongoing.
  The amount of tritium specified was not in the tritiated cathode.
  Looking for new tritium supplier.

## Slide 18: Acknowledgements
Alfredo Galindo-Uribarri
Augusto Macchiavelli
Miguel Madurga – UTK PI
Ingo Wiedenhöver – FSU Lab Director
Brian Schmidt – FSU Sourcerer
Powell Barber – FSU Vacuum Engineer
Jonah Gibbons – FSU Electrical Engineer (Now at Naval Research)
Rick Boisseau and Jason Aragon – FSU Machine Shop
Ben Asher– Previous UTK postdoc
NNSA Grant Number DE-NA0003841
NNSA Grant Number DE-NA0004150

## Slide 19: Thank you!
Questions?

## Slide 20: University of Pennsylvania Tritium Source
Pioneered hydrogen beams from solid-state material
Titanium Sublimation Pumps
Beam current: 0.25 μA – 4 μA
Transmission : ~75% - 33%
Reported 5μC/cm2 in inflector magnet
Middleton, Adams and Kollartis, NIM 151 41-46 (1978)
Notes: Their vacuum system consisted of two Titanium sublimation pumps, one directly after the source and one after the inflector magnet. Titanium sublimation pumps are useful in running tritium beams because they are exhaustless. They operate by running a high current through a titanium filament until the titanium sublimates and condenses onto the wall of the chamber. This titanium will then react with the residual gasses within the chamber thus reducing the pressure. An ion getter pump was also coupled to the system to pump the remaining gasses such as argon.
Titanium sublimation temperature = 1300 deg C
600 hours of operation
Titanium insert wrapped with indium foil (to melt and lock the two pieces together) with an aluminum jacket
The concept of the FSU triton source was derived from the one at the University of Penn

## Slide 21: John D. Fox Laboratory at FSU
TR1: CATRINA Sc.-chamber
TR2:
Resolut SE-SPS Clarion-2
9 MVTandem
Linac

## Slide 22: Pneumatics
Notes: Cathode changes are pneumatically operated, allowing for less breaks in the vacuum

## Slide 23: Power Supplies
Notes: a number of optical and operational elements, namely an Immersion lens, Ionizer, Extractor and Einzel lens for initial beam extraction.
The Cathode, Immersion Lens, and Einzel lens are powered by Spellman SLM10N300 high voltage power supplies while the Extractor is powered by a SLM10P300 power supply. These power supplies are located above the source and controlled by knobs that reach outside the cage
The Boiler is monitored by a thermocouple and heated by a band heater that is connected to 120 V AC and set by a variac controller. The ionizer is powered by DC power supply. Every supply that is located above the source uses the second level platform as local “ground” reference, whose potential is defined relative to the lower platform through the Extractor power supply
The immersion lens serves to focus the cesium ions onto the material to be sputtered; the ionizer is the heating element ionizing the cesium beam; the extractor provides the initial voltage gap for the negative ions to begin their acceleration.

## Slide 24: Cooling
Notes: Coolant for cooling the source and cathodes

## Slide 25: Fume Hood
Florida State University and CENTAUR sponsored duct work installed.
Fume hood designed and built around the source for tritium evacuation
~450CFM  ~100FPM face velocity around the hood
Notes: CFM = cubic feet per minute
FPM = feet per minute
We aim to limit the activity released to the environment from tritium-source operation to less than 200μCi per year

## Slide 26: Fume Hood
Fume hood designed and built around the source for tritium evacuation
Exhaust from roughing pump is released within fume hood
Tritium monitor samples all outgoing air
  Regulatory limit (w/o stack dilution factor): 0.3 Ci /yr
Notes: If any gaseous tritium makes it past the two TSP, the exhaust of the roughing pump is also sent to the fumehood
1 uC/m3
Another monitor at sensitive to 0.1 uC/m3

## Slide 27: Interlock System
Grafana is used to monitor the interlock in the control room
Record what caused a source trip and what level of trip occurred
Record tritium levels, vacuum levels, and TSP1 current

## Slide 28: Deuterium Cathode Fabrication
Titanium absorbs hydrogen  Solid-state hydrogen cathodes
In house deuterium loading, outsourced tritium loading
ARC approved bottle design for their manifold
How much hydrogen can we put into these cathodes?
Courtesy of Dr. Janard

## Slide 29: Deuterium Cathode Fabrication
Insert titanium cathode into bottle and fill with ~500 Torr of deuterium
Heat to >550oC and bake for ~4h
Allow to cool back to room temp. measure gas pressure difference (typically ~200 Torr)

## Slide 30: Deuterium Cathode Testing
Always around a 1:1 ratio of deuterium to titanium determined by the loss of D as well as weight of cathode
This method successfully produced a deuteron beam for experiments within the FSU lab.

## Slide 31: Initial Beam Extraction

## Slide 32: Initial Beam Extraction
~20-40nA used for about 5 days
This consumed ~10% of a single cathode by mass of titanium
1000 hours of beam time per cathode @ 20nA
0.1% of the consumed cathode is ultimately accelerated
Determine where deuterium goes will be where tritium goes
Residual Gas Analyzer showed that very little deuterium makes it out of the source cavity
RGA Mount
~0.25” x 1’ gas restriction tube

## Slide 33: Workshop: Science with Triton Beams
March 14-15, 2024
~50 attendees (in person and virtual) from 25 different institutions
Lab tour
The discussions among peers from various backgrounds allowed for the improvement of the safety features of the triton source and highlighted the importance of the triton source for the advancement of nuclear physics

## Slide 34: Workshop: Science with Triton Beams
Over the course of 2 days there were 15 presentations, i.e.:
  Single-hole states through the Z=50 shell closures through (t,α) reactions
  Charge-exchange reactions through (t,He) reactions
  Triton educed actinide fission and fission dynamics with (t,pf)
  Neutron capture constraints from the Oslo and Surrogate methods through (t,pγ)
  Pairing correlations in nuclei with (t,p)
  (t,p) reactions using the SE-SPS and DAPPER
  Measurement of observables relevant for neutrino-less double beta decay

## Slide 35: Triton beams with light nuclei
Two neutron transfer reactions with light nuclei are well known.
Cross section are large (mb)
We can investigate new physics
Li(t,p)Li

## Slide 36: Triton beams with light nuclei
Li(t,p)Li
  Los Alamos
  15 MeV
  320 μg/cm2 self supporting target
P.G. Young and Richard H. Stokes, PRC 4 1597-1601 (1971)

## Slide 37: Triton beams with light nuclei
Li(t,p)Li
  Los Alamos
  23 MeV
  50 μg/cm2 with carbon backing
F. Ajzenberg-Selove, E. R. Flynn, and Ole Hasen, PRC 17 1283-1293 (1978)

## Slide 38: Triton beams with light nuclei
Li(t,p)Li
  15 MeV
  200 μg/cm2 self supporting target
B = 5 kG
Θlab = 20°
For 10000 counts
  Eex. = 0 MeV (B = 6 kG) ~1.3 hours
  Eex. = 4.30 MeV (B = 5 kG) ~1.1 hours
  Eex. = 6.40 MeV (B = 3.5 kG) ~10 hours

## Slide 39: Triton beams for Actinide structure
U(p,t)U
  16 MeV
  100 μg/cm2
  SE-SPS + Internal Conversion Electron sPectrometer in Coincidence Experiments (ICESPICE)
  Builds the groundwork to study U(n,f) indirectly through U(t,pf)U
U(t,p)U
  15 MeV
  SE-SPS
PI: Mark Spieker
Student: Alex Conley
Bi

## Slide 40: Campaign
3000 hours over two years (assuming half the tritium is usable)
~20 nA estimated from observed deuterium consumption
17 MeV with Tandem, year 1
24 MeV with LINAC, year 2
A committee will be developed to determine the priority of proposals
At end of campaign:
  Remove tritiated cathodes
  Remove TSP1
  Remove “unique” equipment
  Radiation survey

## Slide 41: First Experiment

## Slide 42: First Experiment
L.A. Riley, et al., PRC 103 064309 (2021)
Ti(d,p)Ti
  4820 keV
  4882 keV
  5001 keV
  5427 keV
Ti(t,p)Ti
  J = 5/2- or 7/2-

## Slide 43: 20 Ci of tritium