## Slide 1
S. Krave
General thoughts and notes…
3 December 2025
General Purpose Low Temperature Pressure Prob

## Slide 2: Motivation
Quality Models require validation and good reference data.
  Much can be simulated via multilevel modeling with good results, perhaps better than can be measured but properties assumptions in models should be correct.
Our Universal Test Machines are less than universal
  And/or out of service
Lots of questions on mechanical performance of different conductors
  Delamination performance of REBCO
  REBCO cable properties and limits
  Under-characterized materials
    Telene, 2212, etc

## Slide 3: Infrastructure (FNAL)
At FNAL, for measuring material properties (squishing things) we have the following:
  Low Temperature Loader (90 kN?) (Needs recommissioning)
  Large Instron (500 kN) – out of service (Needs new controller)
  Transverse pressure probe (20 tons)
    Unique specific purpose device
  Small Instron (30 kN)
  Generic 20 ton shop press
    No load cell

## Slide 4: Transverse Pressure Probe
Our existing load probe
Works(ish) down to 4.2K
Low current leads
Not modular
Low precision at low forces
  Force readout is only from non-calibrated hydraulic pressure transducer
  My guess is ~500 lbs hysteresis over 20 tons
    Large errors under 1000 pump psi
    Plus spring return forces
Right now current is being supplied independent of probe (Nitrogen only)

## Slide 5: Proposal: Build a cryogenic capable load system
That happens to have current leads
That happens to fit in a magnet
System is a large pushrod
  Prevents buckling
  Use roller screw or hydraulics to apply either tension or compression to sample
  Use inner and outer supports as current leads
    Wrap in HTS to reduce He Consumption
Where applicable, insert G-10 electrical and thermal brakes
Load Cell
~20 Ton
Load frame to react forces
Double acting cylinder Roller Screw? to apply ~200 MPa to 20mm x 50 mm (20 ton)
Teslatron or generic mount flange
140 mm outer diameter (or VTI OD)
ID Pushrod sized to optimize loads and minimize total heat leak/mass
Instrument thin wall with strain gages and calibrate for force sensing
Interchangeable load head using Instron Style coupling
Some displacement sensing
  Possibly Fiber-optic interferometry for displacement and dilatometry (S. Holl, https://cds.cern.ch/record/2845427/files/CERN-THESIS-2022-274.pdf)

## Slide 6: Mini Transverse Sample Holder : A small step
A smaller low cost holder is desired for testing of larger bend radius ~25mm with straight section samples. It is to work in LN2 without background field. External actuation is desired (~24”-36”)
Max load 5000 lbf
  Limited by load cell
External actuation
Sample width up to 1.5”
Sample Length up to 4”
2 kA (800 stated)
  @10A/mm2 requires 200 mm2 of conductor
  1” x 0.5” copper is 322 mm2 or 6.2 A/mm^2
External Diameter 5.5 inches (fits in 14.7 cm 8.5T magnet)
This would be a little larger in size than the TPI probe.
Actuated by hand pump? – Compression only
In process of lead analysis…
To have in hand:
$15k M&S
.05 FTE
A couple months for simple machining

## Slide 7: Mini Probe Assembled
Right now we can*:
  Cool down in nitrogen tub
  Control to precise set force manually
  Read load
  *Hydraulic fittings have not been installed but it may be directly connected to manual system

## Slide 8: Loadcell Test
Loadcell on device is working well. Set to log to file. Can use for postprocessing.
High precision calibrated loadcell (<0.05% full scale error)
  Can interface with Labview via serial (USB)

## Slide 9
Should Update to IB3

## Slide 10: TS3 Version Extensions
The system will allow loading of a sample in LHe at Test Stand 3 / elsewhere
  The mini-version (small Scale MConE) is extendable to fit into the 8.5T solenoid.
    Extend support rods
    Modify for smaller current leads
The system is planned to allow some substantial current down to 4.2 K (~10 kA) in the 14cm size
It is being configured to be small enough to fit in a background field, to be determined
  8.5T solenoid.
  SMCT?
  HFVMTF

## Slide 11: Controls for Such a system
Initial version can
  Apply load to cold sample, which can be cooled if desired, maybe controlled by hand
A simple mechanical system like this can give simple measurements, but with several adjustable parameters, and generic end implements, it can become incredibly powerful for learning about magnet behavior given some automation
We can add variables to fully realize as MConE
  Parametric load control
  Parametric Current Control
  Parametric field control
  With a VTI… Parametric Temperature control
I will be looking for some help here as things progress

## Slide 12: Conclusions (ish)
A small scale MConE is ready to test, and may be extended for background field testing
  (LN2 only)
In a ‘Full Size’ version, it appears that up to 20 tons and up to 10 kA can be delivered to a sample with a 140 mm diameter probe
  Subject to psu limits
With optimization, it should be reasonable to deliver substantially higher forces and currents.
The system will be compatible with the 8.5T 147mm bore solenoid. This system should be useful for others as well.
Next Steps: Start more detailed design of full size model with plans to scale/adapt for different background fields as they become available. Iterations on detailed thermal analysis before design detailing note:
Questions: Is there anything else to investigate integrating?

## Slide 13: Extra Stuff

## Slide 14: Quick Analysis for ballparking…
Check system resistivity vs diameters
  Using RT properties
Check loads and stresses
  And buckling, note that coaxial configuration eliminates buckling, but probe may move a little
Heat loss assuming a large bar of material with 1 end at 4k and one at room temperature
  Poor assumptions lead to very high heat loads
    Assume 140 mm OD, and vary Inner member
  Since buckling is eliminated, equal area should be optimum.
Targets:
  200 MPa in 1 in^2 sample (29000 lbf)
  2+ kA
  Fits in magnet

## Slide 15
Unsupported buckling limit of ~20 kN
  Again coaxial arrangement eliminates failure by buckling
Column stress of a low value.

## Slide 16: Approximate Sizing for TS3
Heat loads to helium can be calculated as a function of current and shape factor. As a result of the Wiedemann–Franz law, the minimum heat load tends to match across materials, with a varying cross-sectional area
  L = 1.5 m
  Dia = 140 mm (147 nominal)
  Fill factor = .8
  Maximum lead area = .0062 m^2
  10 kA
  IL/A = 2.4E6 (24/10^4 per chart)
    ~ 1W/kA@4.2
    ~ 10 watts to 4.2K at 10 kA
    Optimum for SS is 1700 A
Stainless will work with a large aperture at large currents, but maybe we want to use something more conductive (but less than copper) –need optimization. (6061 Aluminum looks promising)
From A. Ballarino, “Current Leads, Links and Buses

## Slide 17: Clearly…
Detailed analysis is required after some settling
  Include non-linears.
Heat loads from lead power dissipation is substantial and washes out thermal contributions
  High conductivity is better than the alternative
  Thermally insulate device as practical

## Slide 18: Available Facilities: Do we have a preferred initial target?
  64 mm bore 15 T
    49 with VTI
    2400 A for insert
  77 mm bore 14 T
    1850 A for insert
  147 mm bore 8.5 T
    2400 A for insert
  Various cryostats

## Slide 19: Mini Probe Plans
Add electronic proportional valve to programmatically set pressure
  Control via labview
  Integrate with test stand
Automate for multifunction testing

## Slide 20
Teslatron 1 (T1) Cryostat with 15T/17T 64 mm cold aperture solenoid, 2000 A Power Supply w/ 49 mm Variable Temperature Insert (VTI) for testing between 1.5 K to 300 K
Teslatron 2 (T2) Cryostat with 14T/16T and 77 mm cold aperture solenoid, 1850 A PS w/VTI
Teslatron 4 (T4) Cryostat with 8.5T/10T and 147 mm cold aperture solenoid, 2400 A PS
Telslatron 3 (T3) Cryostat with 253 mm neck, 2400 A PS
Physical Property Measurement System Quantum Design PPMS®
4 tube furnaces 6’’ diameter for heat treatment in Argon and Oxygen to 1500C
I 1100°C Three Zone High Pressure Tube Furnace with Gas Flow System
2 m long furnace up to 1250C for heat treatment in air and in argon
Metrology well calibrator for thermocouples calibration up to 700C
Motorized flat-rolling system for wires
Probes and sample holders for SC strand critical current tests up to 2000 A in T1 and T2
Devoted probes to test RRR of up to 6 wire samples in T1,T2 and T3
Sample holders to measure angular dependence of current with field for HTS wires in T1 and T2
Balanced coil magnetometer to measure magnetization in T1 and T2
Device to test critical current sensitivity of impregnated cables to uniaxial transverse pressure in T2
Walters’ spring-type device for tensile/compressive strain sensitivity studies of critical current density in T1
SC transformer for Rutherford cable tests and splice studies up to 28 kA in self-field in T1 and T2
14T/16T Rutherford cable test facility for cable tests up to 30 kA in T2
Modular Insert Test fixture for testing HTS YBCO pancake coils in T2
Hall probe system for magnetic measurements in T4, eg shielding capabilities of bulk MgB2 tubes
Instrons for mechanical properties
Optical and electron microscopes
Respooler and 42-spool cabling machine with turkheads for 1-step and 2-step cables
Superconductor Equipment List at FNAL, IB3A

## Slide 21: IB3A Superconducting R&D Lab
Five ovens up to 1250◦C for heat treatment in Argon and in Oxygen.
Four magnetic cryostats with up to 15T/17 T background field, and with cold apertures between 64 mm and 147 mm are connected to new vent and vacuum systems.
Each system has its own DAQ crate and power supplies up to 2400 A. Variable Temperature Inserts allow measurements between 1.5 K and 60 K.
Is located in IB3-A, a ~6000 square feet addition to Technical Division’s IB3 that was built in 2010 with ARRA funds.

## Slide 22: Experimental Setups
1998: Pressure contact probe for Ic measurements up to 1400 A
1999: Balanced coil magnetometer for magnetization measurements
2001-03: ~28 kA SC transformer for SC splice resistance measurements
2003-2009: Five low resistance probes for stable current measurements up to 2000 A between 1.9 K and 4.5 K
2007: Probe to measure Ic dependence on field orientation for anisotropic HTS wires.
2009-13: SC transformer for cable tests in fields up to 14T/16T and ~30 kA
2001: Device to test Ic sensitivity to uniaxial transverse pressure up to 200 MPa
B

## Slide 23
Modular Insert Test Fixture (ITF) designed to test YBCO double pancake coils and Bi-2212 wire coils
Walters’ Spring probe for strain sensitivity studies of Ic in SC wires
Experimental Setups (cont.)
New keystoned turkhead for Rutherford cable fabrication in 1 step
New RRR probe for wire and bulk samples
New Hall probe system to measure shielding of bulk MgB2 tubes

## Slide 24: Mini probe BOM
Center Pushrod: 1.25” Stainless Steel
Threaded side rods 5/8-11 18-8
  https://www.mcmaster.com/94400A530/
  ~16.9 kip each
Nuts: Silicon bronze:
  https://www.mcmaster.com/93439A665/
  Or Black Oxide: https://www.mcmaster.com/97149A300/
Springs, 5/8 ID die spring 836 LB max each (1600 lbf):
  https://www.mcmaster.com/9588K21/
Load Cell (with USB Computer Logging), to 5000 lbs
  https://www.mcmaster.com/3487N24/
Cylinder: https://www.mcmaster.com/6297K17/
  8830 lbf
Some other parts not yet listed
  Top Cylinder Plate
  Bottom plate
  Mid-braces

## Slide 25
1kW lead power at 685 A for equal areas
~1.7 V

## Slide 26: Bronze
1 kW power at 1635 A for equal cross section
  1V

## Slide 27: Copper
1st pass
  Did not use nonlinear data
  Did not use vapor cooling
  At equal areas, roughly 50 µΩ
    Add HTS because we can.
    Will be lower from RRR affects as well.
  For 1 kW of lead power, this is 4400 A
    .22 V