## Slide 1
Stoyan Stoynev
US Magnet Development Program
Fermi National Accelerator Laboratory
MDPCT1b ( “15 T”): test plan and status
MDP meeting
March 18, 2020

## Slide 2: Magnet test – main goals reminder
Study the effect of the increased pre-load and altered end-support on magnet training and training memory
Measure ramp rate and temperature dependences of magnet quench current
Measure field quality and effects from shims, more detailed studies on decay and snap-back effects
Check safety operation at the high field (stress, MIITs)
Study protection heater parameters (outer layer and middle layer)

## Slide 3: Magnet Test Plan Reminder
Verify that all configuration files are uploaded to the production area and DAQ computers are rebooted
Room Temperature preparation
  Electrical checkouts and Hi-pot
  Hi-pot schedule: Coil to ground (with heaters grounded or floating) at 1000 V, PH to ground (with coils grounded or floating) at 1000 V.
  Check the quench and slow scan data logging
  Verify that the VMTF Overview Display is up and running
  Initial RRR measurements: check CVT, FVT and Lead signals
  Verify that all warm checkouts are completed
  in the Test Readiness Verification Procedure document (TID-N-769)
  Magnetic measurements (z-scan) in VMTF at room temperature
Cool down
  Restricted cool down:
  temperature difference between the magnet top and bottom should not exceed 50 K
  Periodically check the SG data logging
System checks at 4.5 K with 60 mΩ dump resistor:
  Cold Electrical Checkout and Hi-pot, Quench Detection Checkout
  Verify that all cold verification steps are completed in the Test Readiness Verification Procedure (TID-N-769)
  Connect the CPS3 power leads and proceed with the AQD/DQD balancing
  Ramp to 1000 A, adjust Cu-I AQD balance and perform manual trip
  Heater induced quenches at up to 5 kA
We are here

## Slide 4: Magnet Test Plan Reminder (2)
At 1.9 K operation with 60 mΩ dump resistor:
Cold measurements after training may have to become part of the second thermal cycle!
It depends on test timing and schedule optimization.
  The maximum current will be limited to 11 kA
  Magnetic measurements at 5.0 kA– see the Magnetic measurements plan
    Z-scan of the main field and field harmonics with both probes
    Decay and Snap-back measurement with both probes at 760 A during current ramp up
  Quench Training at 20/50 A/s to a plateau
    Plateau is 5 quenches with no improvement above 50 A
    More than one detraining (above 200 A) quench or multiple detraining quenches at above 50 A will urge a team review meeting
    Flux-jump measurements
    Strain gauge and acoustic signal measurement
  Ramp rate studies (10/100/200/300/400 A/s)
  Magnetic measurements in cycles up to Imax - 500 A – see the Magnetic measurements plan
    Z-scan at different currents (1/5/9 kA) with both probes
    Decay and Snap-back studies with both probes at currents 300 A and 750 A during current ramp up and down
  Temperature dependence studies (2.1/2.7/3.4/4.0/4.5 K)
At 4.5 K operation with 60 mΩ dump resistor:
  Inductance measurements at 50 and 300/200 A/s
(and up to 80% of the current reached in training)
  Splice resistance measurements

## Slide 5: Magnet Test Plan Reminder (3)
Warm up to 300 K
  Cold RRR Measurements.
  Periodically check the SG data logging
At 300 K after the test
  Repeat RRR measurements at room temperature
  Magnetic measurements (if needed)
Cool down to 1.9 K (second thermal cycle)
Cold measurements from the previous thermal cycle after training may have to become
part of this one (see above)! Then they will be preformed before the heater tests.
  Check training memory
  Training to plateau if necessary
  Protection heater studies (if safe)
    minimum power to quench vs current
    time delay vs current
    quench distribution and propagation (L2 PH are in between coils)
At 4.5 K operation
  2 training quenches at 20/50 A/s
Warm up to 300 K

## Slide 6: Latest check-out (in-dewar)
All VT detection and characterization channels were fine except:
C3
C4
C7
C5
C6
D3
D4
D7
D6
D5
D1
D2
C1
C2
Layer 3
Layer 4
Coil 5, VT D3 is N/A –
not new, known issue, can’t be fixed;
VT to be skipped as before
Coil 4, VT C5 is flaky (sometimes N/A) –
[V]
[V]
It suggests the VT “insert” is touching the next turn.
Coil 4, VT C5 is to be skipped in the readout.
Coil 4, VS C5_C4
From warm
RRR data :
Coil 4, VS C6_C4

## Slide 7: Latest check-out (in-dewar) (2)
Hipot (1000 V) passed, but some readings were higher than usual (still within limits)
Coil-to-ground (heaters floating): 6.2 A (non-linear dependence with voltage)
The coil-to-ground resistance was measured (with ohm-meter) to be below 2 G
Coil-to-ground (heaters grounded): 4.8 A (non-linear dependence with voltage)
Heater-to-ground readings were below 0.2 A
Earlier hipots on the magnet itself (no top header) showed < 0.2 A currents for all

## Slide 8: Latest check-out (top header)
RTD: One of three temperature sensors on the shell failed – won’t be repaired (it is at the LE)
However, a repaired connection brought the bore temperature sensors back to two
QA: Flexible quench antenna (on the bore tube) channels – fine.
LBNL QA (external) to be deployed later
Acoustics: Four acoustic sensors (“old” style) operational.
(we constantly fail to get comparable results with the “new” type of acoustic devices;
we strongly think sensors should be bolted and not glued to surfaces)
SG: All Strain gauges are operational.
LE1
LE2
RE1
RE2
Voltage readings (different scales) from acoustic sensors testing – magnet hanging over/in the pit

## Slide 9: Warm RRR data and magnetic measurements
All CVT and FVT channels fine, some examples below
Warm magnetic measurements ongoing…
Current (10 A)
Whole/half coil signals

## Slide 10: Next steps
Restricted T<50 K cool-down over the weekend, possibly starting tomorrow
last time it took 76 hours (~ 3 days)
Cold check-outs on Monday
will likely take two days
Training as soon as Wednesday
It is reasonable to expect that there will be less people available for support in next weeks
and priorities shifted resulting in delays if not suspensions.
However:
We may be allowed to run the test within the minimal Lab operations being prepared
(and eventually invoked)

## Slide 11: Summary
Following the developed run plan
No signifficant issues encountered, “warm” steps being finalized
Planning for
- restricted T cool-down over the weekend (starting tomorrow?),
- cold check-outs on Monday and
- training later that week
We may be allowed to run the test within the minimal Lab operations being prepared
(and eventually invoked)

## Slide 12: Back up