## Slide 1
X. Wang
REBCO WGM #7, 6/18/2020
C2 magnet: fabrication and test results

## Slide 2: C2 aims to generate 3 T with longer conductors, metal mandrels and Stycast to constrain the conductors
Parameters | value
Tilt angle, L1/2 | 50°
Tilt angle, L3/4 | 35°
Min. bending radius | 30 mm
Transfer function | 0.47 T/kA
Target dipole field | 3 T
Layer 1
2
3
4

## Slide 3: C2 used the state-of-art 30-tape CORC® wire to boost the magnet current
Wire ID | Length (m) | Wire OD (mm) | Average tape Ic (A) 77 K, SF | Peak field on wire (T) | Min bend radius (mm)
C2-L1 | 18 | 3.80 | 70 | 3.6 | 30
C2-L2 | 20 | 3.80 | 70 | 3.6 | 35
C2-L3 | 24 | 3.77 | 69 | 3.2 | 30
C2-L4 | 28 | 3.67 | 57 | 2.8 | 35
SuperPower tapes: 2 mm wide, 30 µm substrate, 5 µm surrounded Cu stabilizer
Layer 4 wire contains high- and low-pinning tapes – conductor grading
Measurements at ASC/FSU detected deviation from expected tape performance, allowing the quick feedback to SuperPower and wire design at ACT

## Slide 4: We expect C2 to reach 3 T at 4.2 K based on the in-field performance data of a Layer 1 sample
V(I) transition of a Layer 1 sample measured by ACT at different background fields
Bending radius close to C2
Jeremy Weiss and Danko van der Laan

## Slide 5: Mechanical analyses confirm that C2 and beyond needs a metal mandrel
C2 mandrel used Aluminum Bronze to leverage the experience of the CCT program at LBNL [D. Arbelaez, Thur-Af-Or24-05]
Mandrel stresses of 150-200 MPa is too high for 3D printed Bluestone that was used for C1 mandrel
Lucas Brouwer
Aluminum
Aluminum Bronze
Bluestone
Notes: - Cooldown to 4.5 K minus Lorentz Force

## Slide 6: It may not be straightforward to machine the tilted groove in a metal mandrel
We tried with an in-house 5-axis CNC machine
Mandrel length is limited to about 300 mm

## Slide 7: We reduced groove depth on mandrels as an interim solution to continue the magnet development
Laura Garcia Fajardo
Bill Ghiorso, Hugh Higley
Meanwhile, we used half-depth groove to help with winding
Full-depth radial groove: Straightforward to machine but challenge to wind with tension
We will continue investigating this

## Slide 8: We chose the half-depth groove design for C2 magnet
Tilted
Full-depth radial groove
Half-depth radial groove

## Slide 9: So we started
Winding with Hugh’s winding table
Mandrel designed by Bill Ghiorso and machined by Maxwell Maruszewski

## Slide 10: Co-wound voltage-tap wires with the conductor to reduce the inductive pickup during tests
Voltage tape wires

## Slide 11: Painted Stycast 2850MT after winding to constrain the conductors
Wrapped fiber cloth on top of conductor
Wrapped and heated heat shrink tapes around coil before Stycast cured
Hugh Higley

## Slide 12: We measured the transport performance of each layer at 77 K before making the next layer
Hugh Higley and Timothy Bogdanof

## Slide 13: Few months layer, four layers of C2 oriented for a dipole configuration
Layer 4
2
1
Terminals
Cu blocks covering terminals
Aluminum Bronze mandrel
3

## Slide 14: Assembly and insert into the cryostat

## Slide 16: We used a systematic and progressive approach to develop and understand the fabrication technology for C2
3-turn
40-turn
Wind
Test at 77 K
Apply Stycast
Test at 77 K
Test at 77 K
Assembly
Test at 77 K
Test at 4.2 K
Wind
Test at 77 K
Apply Stycast
Test at 77 K
Layer 1
Layers 2 – 4
Wind
Apply Stycast
Test at 77 K
Assembly
Test at 77 K
Test at 4.2 K

## Slide 17: Layer 4 showed lower performance than the other three layers at 77 K but get closer to Layer 1 after assembly
~ 30% Ic reduction due to higher fields
There is potential to perform at 4.2 K as the voltages across Layers 1 and 2 start rising

## Slide 18: Layers became superconducting in sequence, indicating cooling from outer to inner layers
Layer 4
3
2
1
Cool down from 100 to ~ 90 K

## Slide 19: A controlled increase in the maximum current allowed us to probe the true performance of C2
Ramp 4
Increasing the threshold for quench detection

## Slide 20: A controlled increase in the maximum current allowed us to probe the true performance of C2
Ramp 11
Increasing the threshold for quench detection

## Slide 21: A controlled increase in the maximum current allowed us to probe the true performance of C2
Ramp 13
Increasing the threshold for quench detection

## Slide 22: A controlled increase in the maximum current allowed us to probe the true performance of C2
Ramp 15
Increasing the threshold for quench detection
Reproducible V(I) transition between ramps
  n = 13.1 for Layer 4
  n = 6.8 for Layer 1, consistent with the behavior of short sample measured at ACT
Wire Je = 595 A/mm2 at 6.3 kA

## Slide 23: The current-carrying capability increases by a factor of ~ 10 from 77 to 4.2 K

## Slide 24: Layer 1 Ic degraded by 5% at 4.2 K after the thermal runaway

## Slide 25: Inter-layer joint resistance ranges from 8 to 24 nΩ at 4.2 K - room for improvement
Joint voltage is measured across the voltage taps in each layer
Layer 1/2
Layer 3/4
2/3

## Slide 26: We measured the field quality of C2 and observed strong persistent-current effects at the magnet center
Measurement with a 100 mm long rotating coil developed by J. DiMarco at FNAL
B1
b3

## Slide 27: Stronger effects at 4.2 K – 30 tapes with large Jc
Cory Myers from OSU will present detailed field quality study, Thu-Mo-Po4.07-05
b3
B1

## Slide 28: A ramp-rate dependence was observed in non-allowed skew terms. A similar behavior was seen in C1 with the plastic printed mandrel
C2 on aluminum bronze mandrel
C1 on bluestone mandrel
Inter-tape coupling currents?

## Slide 29: We close the chapter on C2 with a good result with more questions arise
C2 reached 2.9 T, 98% of the expected value
Layer 1 started transitioning at 4.8 kA, 73% of the short-sample prediction. Why?
What caused the low performance of Layer 4?
Where is the heat/voltage generated?
How can we improve for the next magnet?
“A good result and more questions. That’s what we need.” – S. Gourlay