## Slide 1: Bi-2212 Rutherford Cable Solenoid: “Baby-Ruth” Update
Daniel S. Davis, Ulf P. Trociewitz, George Miller, Youngjae Kim, Ernesto S. Bosque, Jozef Kvitkovic

## Slide 2: Reinforced Baby-Ruth Coils: +2 T in 14 T LTS
Coil Specs:
  6.45m conductor
  45 mm ID, 6 layers x 6 turns
  2.781 kA (0.75*Ic) at 16.64 T load-line,
  Self Field (1.97 T central field, 2.64 T peak on conductor)
Mechanics:
  At 3.7 kA, 14 T
  400 MPa source stress
  0.38 % strain
  --> we should still have some stress margin left
Btot in coil, Iop=0.75% IC

## Slide 3: First test coil “Baby-Ruth” Fabrication Well Underway
5 kA leads
  7.5 kA VCL installed on probe head
  5 kA AMI LTS received
Probe supports assembled
Coil terminals
  Ag terminals fabricated
  Cu terminal and soldering fixture received
14 T LTS
+2 T HTS

## Slide 4: Baby-Ruth Wound Ready for OPHT
9-strand cable handles very well during coil winding
Coil has passed through the organics burn-out heat treatment
Awaiting repairs of the Deltech furnace for 50-bar OPHT
Test must wait for helium plant start-up in new year
Individually sealed wire ends
Baby-Ruth As-Wound Test-Coil

## Slide 5: Jogged Solenoid Winding With 9-strand Cable
Ag Voltage Flags
Jogged (un-pitched) windings
Underfilled layer transitions
Smooth hard-way bends into terminal

## Slide 6: Map for Rutherford Cable Solenoids from MDP Meeting
2020 Q1/2 | 2020 Q3/4 | 2021 Q1/2 | 2021 Q3/4 | 2022 Q1/2 | 2022 Q3/4
Coil-1 Engineering Design
Coil-1 Fabrication
Coil-1 Test
Coil-2 Engineering Design
Coil-2 Fabrication
Coil-2 Test
Coil-2 Cable Fabrication
Coil-1 Review
Coil-2 Cable Determination
Coil-2 Review
Coil-3
Coil-3 Cable Determination
Coil scale-up options contingent upon successful tests
2-3 T added (17 T total), 8 meter (existing cable) = “Baby-Ruth”
5-6 T added (20 T total), 40 meter needed +back-up
Possible Full-Scale: 100’s of meters for full volume (11-18 T added 25-30 T total). Major funding and large furnace required.
Testing in 14 T B.G.
6
Fabrication slipped due to unforeseen resource limitations and COVID delays

## Slide 7: Renegade Furnace
Deltech
Had a zone failure that still needs to be addresses
Renegade activities currently have higher priority
Renegade
New heater insert is a significant better fit than the previous one
Heater insert vacuum formed in one piece instead of stacked rings
Actual bore now 279 mm
Insert and pressure vessel have now been assembled and feedthroughs are being installed
Crane service to come in this next Thursday to install assembled pressure vessel into hydraulic stand and install inside room
Thermocouples, el supply lines from control panel, and some of the gas management system to be installed by end of next week early Jan
Pressure and power tests (in that order) expected by Jan

## Slide 8: 2212 Coating
Rutherford cable coating
Experimented with two coating methods, airbrushing and hand painting
  Airbrushing: fairly homogeneous coating possible but generates lot of waste due to overspray as well as spider-webbing, which could not be reduced significantly by changing amount and ratios of the solvents (ethanol, xylene)
  Hand painting: produces significantly less waste but coating of varying thickness, second coat can remove first one
  Both methods currently lack heated drying to allow application of polymer top coat
A mechanized coating facility appears desirable for both, accelerator coil and solenoid, effort
A cable coating facility is likely to look substantially different from the existing strand coating facility since a mechanized brushing or rolling application of the slurry may require a horizontal layout
With an increase in the lengths of to be coated cables it appears to be worthwhile to start thinking about setting up such a facility

## Slide 9: Conductor Insulation Options
Spray
  Wasteful of coating
  Spider-webbing can lead to agglomeration on surface
  Limited solvent choices(ethanol, xylene)
Hand coat
  Labor intensive for even such a small test coil
  Single coat limitation, varied thickness
Mechanized roller/brush system
  How uniform would the application be
    Keeping applicator clean of buildup
    Moving applicator or cable?
Dip-coat
  Would need to re-tool entire existing system
  Moving seal for rectangular cross section
    Continually saturated paint roller or sponges
    How to keep from gumming up?
    Alternative horizontal angled bath entry/exit
  Uniformity of edge coating?
Die coat
  Industrial operation, requires high speed -> furnace heat upgrades